CNSE NanoFab Tool Calibration Form

DATE: __2/18/2016________________

Tool Description: ____Unaxis 790 PECVD__

Test: SiN deposition rate

Run Parameters:

Sample size: 4” silicon wafer

Nitride Recipe:  Standard

Deposition time:  10 min

Deposition temp:  250 C

Process Gases:  SiH4, N2, NH3

Average Thickness (9Pts), 10 mm edge exclusion:  1032 A

Range: 28 A

Standard Deviation:  9

Uniformity:  1.3%

Deposition rate:  103 A / min

Engineer: CNSE Staff_Frank Lee_________________