CNSE NanoFab Tool Calibration Form

DATE: __2/18/2016________________

Tool Description: ____Unaxis 790 PECVD__

Test: SiO2 deposition rate

Run Parameters:

Sample size: 4” silicon wafer

Oxide Recipe:  Standard

Deposition Time: 10 min

Deposition Temp: 250 C

Process Gases:  SiH4, N2O

Average Thickness (9Pts), 10 mm edge exclusion:  5669 A

Range:  183 A

Standard Deviation:  59

Uniformity:  1.6%

Deposition rate:  567 A / min

Engineer: CNSE Staff_Frank Lee_________________