CNSE NanoFab Tool Calibration Form

DATE: __6/20/2016________________

Tool Description: _E-beam Evaporator #2_______

Test: Deposition thickness and uniformity

Run Parameters:

Sample size: 4” silicon wafer

Material: Titanium 100 ang., Palladium 200 ang.

Starting chamber pressure: 4.0 x 10-6 Torr

Rotation speed:  50%

Evaporation rate: 1 ang/sec.

Tooling factor: 84%

Calibration tool: Dimension 3100 AFM

Results:

Uniformity:  1.3% using three measurements (1 top, 1 center, 1 bottom)

Requested thickness: 300 angstrom

Measured thickness: 304 angstrom avg.

Engineer: CNSE Staff_Frank Lee_________________