CNSE NanoFab Tool Calibration Form

DATE: __2/23/2016________________

Tool Description: ____STS Multiflex RIE__

Test: PECVD silicon dioxide etch rate

Run Parameters:

Sample size: 4” silicon wafer

Mask:  Shipley 1813 final spin at 4K rpm

Edge bead removal

Etch Condition:  100 mT chamber pressure

CHF3: 40 sccm

CF4: 10 sccm

300 W RF

PECVD silicon dioxide etch rate:   219 nm / min

Etch Profile:  80 degrees (estimated)

Oxide to photo resist selectivity:  5:1

Etch Uniformity:  1.3 %

Engineer: CNSE Staff_Frank Lee_________________