Oxford Cobra Silicon Trench Etching System

APPENDIX A

Abridged Tool Operation Instructions

- Pre-etch Chamber Clean

  - Check the load lock. Make sure the Si clean/condition wafer inside is clean and correctly positioned.

  - From the “System Pumping Screen,” use the “Process Menu” to navigate to the “Recipes” page.

  - Load the appropriate chamber clean recipe for your intended etch process.

  - Edit the “Step Time” of the chamber clean to 10 minutes. DO NOT CHANGE THE OTHER PARAMETERS!

  - Press Run to run the Pre-etch Chamber Clean.

  - After the process completes, accept the YELLOW ALERT and wait for it to return to the “System Pumping Screen.”

- Pre-etch Chamber Condition

  - From the “System Pumping Screen,” use the “Process Menu” to navigate to the “Recipes” page.

  - Load your desired etch process.

  - Edit the “Step Time” of the etch step to 10 minutes. DO NOT CHANGE THE OTHER PARAMETERS!

  - Use the “Process Menu” to navigate to Chamber 1 and wait for the tool to reach temperature.

  - Use the “Process Menu” to navigate back to the “Recipes” page.

  - Press Run to run the Chamber Condition.

  - After the process completes, accept the YELLOW ALERT and wait for it to return to the “System Pumping Screen.”

- Wafer Load

  - On the “System Pumping Screen,” press STOP under “Cycling Loadlock Pumping.”

  - Wait a few seconds and then press VENT under “Cycling Loadlock Pumping.”

  - Wait until the “Vent Time Left” is under 90 seconds before opening the load lock.

  - Replace the Si clean/condition wafer with your own. Make sure to position it properly

  - Close the load lock door and press STOP under “Cycling Loadlock Pumping.”

  - Wait a few seconds and then press EVACUATE under “Cycling Loadlock Pumping.”

- User Etch

  - From the “System Pumping Screen,” use the “Process Menu” to navigate to the “Recipes” page.

  - Load your desired etch process.

  - Edit the “Step Time” of the etch step to your desired time. DO NOT CHANGE THE OTHER PARAMETERS!

  - Press Run to run the etch process.

  - After the process completes, accept the YELLOW ALERT and wait for it to return to the “System Pumping Screen.”

- Wafer Unload

  - On the “System Pumping Screen,” press STOP under “Cycling Loadlock Pumping.”

  - Wait a few seconds and then press VENT under “Cycling Loadlock Pumping.”

  - Wait until the “Vent Time Left” is under 90 seconds before opening the load lock.

  - Replace your wafer with a Si clean/condition wafer. Make sure to position it properly

  - Close the load lock door and press STOP under “Cycling Loadlock Pumping.”

  - Wait a few seconds and then press EVACUATE under “Cycling Loadlock Pumping.”

- Post-etch Chamber Clean

  - From the “System Pumping Screen,” use the “Process Menu” to navigate to the “Recipes” page.

  - Load the appropriate chamber clean recipe for your intended etch process.

  - Edit the “Step Time” of the chamber clean to the same amount as your etch time, or to 5 minutes if your etch was less than that. DO NOT CHANGE THE OTHER PARAMETERS!

  - Press Run to run the Post-etch Chamber Clean.

  - After the process completes, accept the YELLOW ALERT and wait for it to return to the “System Pumping Screen.”

  - Fill out the log book

  - Close the valve on the liquid nitrogen tank if necessary

  - Clean up after yourself and make sure everything is back where it belongs before you leave