CNSE NanoFab Tool Calibration Form

DATE: __2/10/2016________________

Tool Description: ____Karl Suss MA-6 Mask Aligner___

Test: S-1813 Resist Exposure/Dev. Critical Dimension

Run Parameters:

Sample size: 4” silicon wafer

Material: S-1813 Photoresist

Spin Speed: 4000 RPM

Bake Temp: 115C on hotplate

Bake Time: 60 sec.

Exposure Tool: Suss MA-6

Exposure Mode: Vacuum Contact

Exposure Time: 7 sec.

Mask: Resolution 1

Developer/Time: AZ 400K 1:4/ 60 sec

Results:

CD: 10 um L/S

CD Measured (4 quadrants average) : 10.0125 um

Engineer: CNSE Staff_Frank Lee_________________