N

E

Quintel Mask Alignment System

Q4000 Series

Introduction

The Quintel mask alignment system is an integrated optical- mechanical, pneumatic- electrical system which allows accurate alignment of sensitized semiconductor wafers or substrates with a mask and exposes them to ultraviolet radiation to transfer the pattern of the mask to the substrate for further processing on the way to producing a semiconductor or other microelectronic devices. Optical and mechanical functions of the aligner are coordinated in a manner that will allow an operator to learn the system in a few minutes. The entire mask alignment procedure is designed for fast, convenient operation.

System Control Features

- Quintel Power Supply 365-405nm source; bottom shelve right side.

- Quintel Light Source A&B; under table.

- Ultrak Monitor; top right.

- MV200 view controller; top right.

- Joy Stick 1, mask viewing manipulator; left side.

- Joy Stick 2, substrate manipulator; right side.

- Theta substrate adjustor; left side.

- Microscope illumination control panel & vacuum/ separation pressure gauge; left side.

- Objective focus switch/ course align/ contact separation button; right side.

- Head Lift Exposure Source with zoom objectives; center.

- Mask Holder; center. (Note. Mask holder is fixed)

- Sample holder vacuum chuck; sliding stage center.

- Split field objectives with zoom capabilities; center. (Note. Zoom course top, zoom fine bottom)

- Parameter Control display panel; right side.

- Power On switch; right side.

- Head Lift lock switch; right side.

- Expose/ mask clamp/ mask vacuum/ clear/ load buttons; right side.

- Objective slide knobs; located on right and left side of objective holder.

Quick and Easy Operation Procedures

- Check if Lamp power supply is on and reading 275watts. (If lamp power supply is off; then push power lever to the on position; then press the start button and wait a few minutes for the lamp power to come up to 275w).

- Now push the “Main Power ON” button.

- Push Head Lock “ON” button.

- Turn “ON” Ultrak monitor.

- Turn “ON” MV200 controller.

- Turn “ON” Objective Light Source A&B.

- Load Mask onto holder. Face Down.

- Press Mask Vacuum Button.

- Push Enter on Parameter Display panel to get “Main Menu”.

- Press # [1] for exposure Parameter; [To set exposure time press 0, enter time and press enter]; then enter again to return to “Main menu”.

- Press # [3] for Op mode.

- The display should show [*] exp pressure; If not Press # 1 for [*] exp pressure and then press Enter twice. Display menu should be in unloading mode. (NOTE. [*] should only be on exp pressure, if others depress)

- Pull sample stage towards you and load sample onto sample chuck. (For small pieces, select correct size plastic vacuum chuck adapter and center on chuck.) *Chuck vacuum is automatic once stage is pushed ¼ of the distance in toward the align position.

- Push stage forward into align position.

- Press “Load” button on; [Chuck Stage will come up into separation mode].

- Separation Mode [make necessary alignments].

- When alignment is complete, Push “Contact” button.

- Press “Expose” button; [Expose button will bring Head Lift forward and open light source shutter] Exposure is complete.

- Unload sample; Pull stage toward you; [repeat steps for next exposure].

- If complete remove mask; press Mask vacuum to release.

- Press Enter for Main Menu.

- Press # [6] for Power Off and Turn Off monitor, controller and light source.

Note. Press clear to abort any step and then enter to return to Main Menu.

*The system is currently configured for fragment pieces and 2    inch wafers.