N

E

RECOMMENDED LITHOGRAPHY PROCEDURE

FOR Microposit S1813 RESIST

POSITIVE TONE RESIST:

Note: There is no magic process that works for all cases. A firm understanding of what you are trying to accomplish will determine the actual process which will work. This recipe should be seen as a starting point for many commonly useful processes.

1. Clean wafers:

A. Remove photoresist if wafers are being reworked using acetone or Microposit Remover 1165. (see step 11).

B. (Depending on previous steps) RCA or sulfuric peroxide cleans are preferred but as a MINIMUM : 5 ' ultrasonic acetone, 5' ultrasonic isoproponol, 5' DI water rinse, Nitrogen blow dry (Do NOT let water evaporate from wafer, rewet if necessary).

2. Dehydration Bake:

A. Hot Plate: 5' @ 150°C minimum (more is better)

OR

B. Oven bake: 20'@250°C minimum (higher temp is better).

3. HMDS MP-P20 Prime:

CAUTION: HMDS IS HIGHLY FLAMMABLE AND POISONOUS. USE ONLY AT FUME HOOD, DO NOT BREATHE VAPORS AND KEEP AWAY FROM METHANOL. READ MSDS SHEETS

Instructions for Use:

- Make up: MicroPrime MP-P20 Primer 80/20 should be used as supplied.

- Procedure/Use: Priming just prior to photoresist coating improves adhesion by producing an interfacial bonding layer for the resist and by removing any moisture from the substrate surface.

Spin priming may be done by puddle. The primer should be in contact with the surface for a minimum of 10 seconds followed by spin drying for 20 - 30 seconds at 3000 - 5000 rpm or until interference fringes are no longer observed.

- Bake at 105 -115C for 1-3 minutes

4. Spin Coat Wafers:

Spin speed depends on desired thickness. 4000 RPM for 30 seconds gives a coating after bake of a little more than 1 micron as shown below. Make sure resist has warmed up to room temperature before applying.

Approximate thickness (per manufacturer) vs spin speed after softbake:

3000 RPM 1.6 µM

4000 1.4

5000 1.25

6000 1.15

7000 1.0

For more exact thickness data prior to develop, you may scratch the resist in an area of non-interest, and use the Dektak 8 AFTER baking PR.

5.Soft Bake:

A. Hot Plate: 1 minute @115°C

OR

B. Oven: 20 minutes @ 95°C

6. Exposure:

Exposure depends on resist thickness, lamp intensity and many other variables, but for 4000 rpm with exposure intensity at 13mW/cm2, (280W, lamp power), 4-6 seconds is a good starting point on the NanoFab MA-6.

7. Develop:

Develop in beaker with 5:1 water: Room Temperature Microposit 351 developer with constant gentle agitation. (You should try to keep this step constant and adjust your exposure in step 6 to get desired results. Also soft baking affects this step greatly. Under-baked resist will develop very fast with little control. Over-baked resist will develop slowly and may not clear at all.) Rinse in flowing DI water immediately for at least 2 minutes and blow dry with nitrogen. (Do not allow water to evaporate).

8. Postbake:

Hotplate: 2 minutes @150°C

OR

Oven: 30 minutes @120°C minimum

Note: This step depends heavily on etch or ion implant step. Higher temperatures help image stability, adhesion and plasma/chemical resistance but above 120°C, stripping PR requires O2 plasma, acid strip or Microposit Remover 1165. Postbake up to 150°C may be done without resist flow.

9. Descum:

It is strongly recommended that an oxygen plasma descum step be used especially for small geometries such as contacts and vias and especially if a dry etch step is being used. The RIE should be used with O2 at a power of maybe 100-150W to remove about 500Å of PR. Use the Filmetrics to verify proper PR removal. If wet etching without this step appears to be taking longer than calculated then a PR scum layer may be interfering with the etch and the addition of this step may be required.

10. Wet Etch:

For Oxide etching it is recommended to use BOE etchant rather than HF-water mixtures. Undercut and resist lifting are much worse for HF-water than for BOE. For very long etches (~hrs) even in BOE, lifting may occur. For these cases, intermittent hard baking, resist hardening (UV or PRIST) or the use of a hard mask may be necessary.

11. PR Stripping:

For photoresist NOT exposed to ion beams, plasmas or temperatures in excess of 120°C, Microposit 1165 or acetone may be used. Otherwise, O2 plasma or an acid strip (Pirhana, Caros acid etc.) may be used (depending on the underlying layer). Wafers must be free from old photoresist and organics before re-coating them.